Acoustical Society of America (ASA) Working Group S3-79

ANSI S3.5-201X: Methods for the calculation of the Speech Intelligibility Index

Determine the Coupler to Free Field Difference according to the ANSI RETSPL

Author: Koenraad S. Rhebergen

K.S.Rhebergen@UMCUtrecht.nl

Draft for extension to ANSI standard S3.5-1997; Section 5 after present sections 5.1.4

5.1.5	Determine the Coupler to Free Field Difference in case of headphone or insert measurements

    - Aim: Determine the Coupler to Free Field Difference.

This section describes the basic principle of calculating the Coupler to Free Field Difference (CFFD). SII model predictions are based on the speech and noise spectrum levels in the free field. Speech intelligibility predictions with headphone or insert can only accurately be modelled by the SII model if the speech and noise spectrum levels are transformed from coupler levels to Free Field levels. The CFFD is dependent on the used transducer (e.g. TDH39, HDA200, ER3A) and the used coupler (e.g. IEC303, IEC318, IEC711). The CFFD for a band is given as:

Qi =  xi – ri		(x)

Where Qi is the CFFD for band i, xi  is the Minimum Audible Field, and ri is the Reference Equivalent Threshold Sound Pressure Levels (according to ANSI S3.6 – 2010 and PTB Test Report Audiometric Headphones HDA 300) for the used transducer calibrated on a specific coupler.

5.1.6	Transform the speech and noise spectrum levels in a coupler to free-field levels

The equivalent speech spectrum level, E’ transformed from coupler to free field is given by:

E’i = Ei + Qi			(xx)

Where Ei is the equivalent speech spectrum level for band i measured in a coupler

The equivalent noise spectrum level, N’ transformed from coupler to free field is given by:

N’i = Ni + Qi			(xxx)

Where Ni is the equivalent noise spectrum level for band i measured in a coupler

Table xxx – One-third octave band SII procedure – frequencies, free-field to eardrum transfer function, Minimum Audible Field (MAF), and free-field to coupler difference.

|  |  |  |  |  |  | Coupler to Free-Field Difference(Q) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | coupler: | IEC711 |  | HA2 | HA1 | IEC 318 | IEC 318 | IEC 318 |  |  |
| Band no. | Frequency (Hz) | Free field to eardrum | MAF | transducer: | ER 3A |  | ER 3A | ER 3A | TDH39 | HDA200 | HDA300 |  |  |
| 1 | 160 | 0 | 18.3 |  | -6.2 |  | -7.7 | -8.2 | -20.2 | -7.7 | -6,7 |  |  |
| 2 | 200 | 0.5 | 15.3 |  | -6.2 |  | -2.7 | -4.2 | -17.2 | -6.7 | -7,5 |  |  |
| 3 | 250 | 1 | 12.7 |  | -4.8 |  | -1.3 | -1.8 | -14.3 | -5.3 | -7,3 |  |  |
| 4 | 315 | 1.4 | 10.5 |  | -5 |  | -1.5 | -4.5 | -11.5 | -5 | -6,4 |  |  |
| 5 | 400 | 1.5 | 8.7 |  | -4.3 |  | -0.3 | -6.3 | -8.3 | -4.8 | -4,1 |  |  |
| 6 | 500 | 1.8 | 7.5 |  | -2 |  | 2 | 1.5 | -6 | -3.5 | -0,5 |  |  |
| 7 | 630 | 2.4 | 6.5 |  | -1 |  | 2.5 | 2 | -4 | -1.5 | -0,2 |  |  |
| 8 | 800 | 3.1 | 5.9 |  | 0.4 |  | 4.4 | 4.4 | -2.6 | -0.1 | 1,1 |  |  |
| 9 | 1000 | 2.6 | 5.7 |  | 0.2 |  | 5.7 | 5.7 | -1.8 | 0.2 | 3,7 |  |  |
| 10 | 1250 | 3 | 5.3 |  | -3.2 |  | 3.3 | 4.3 | -2.2 | -0.7 | 2,8 |  |  |
| 11 | 1600 | 6.1 | 4.1 |  | -5.4 |  | 2.1 | 2.6 | -3.9 | -1.4 | 1,7 |  |  |
| 12 | 2000 | 12 | 2.5 |  | -9 |  | -0.5 | 0 | -6.5 | -2 | 2,5 |  |  |
| 13 | 2500 | 16.8 | 0.3 |  | -13.2 |  | -4.7 | -4.2 | -10.2 | -2.7 | 1,8 |  |  |
| 14 | 3150 | 15 | -1.6 |  | -14.6 |  | -5.6 | -4.1 | -13.1 | -5.6 | 1,0 |  |  |
| 15 | 4000 | 14.3 | -1.9 |  | -16.9 |  | -7.4 | -1.9 | -13.9 | -11.4 | -1,4 |  |  |
| 16 | 5000 | 10.7 | 1.4 |  | -17.1 |  | -3.6 | -0.1 | -9.6 | -12.6 | -9,1 |  |  |
| 17 | 6300 | 6.4 | 10.4 |  | -5.6 |  | 8.4 | 12.4 | -10.6 | -7.1 | -10,9 |  |  |
| 18 | 8000 | 1.8 | 20.6 |  | 5.1 |  | 20.6 | 24.1 | 5.1 | 3.1 | -2,4 |  |  |

Note 1: the CFFD (Q value) for one octave band noise should be based on the energetic means of the 1/3 octave CFFDs surrounding the desired octave band center frequency. However, the CFFD for 8kHz cannot be determined for the TDH39 and ER 3A because there are no official RETSPLs values for 10kHz and thus no CFFD for 10 kHz is available for these earphones.

Note 2: the CFFD (Q value)in this section is based on the MAF (monaural) according to ANSI S3.5. Therefore the CFFD is only suitable for SII calculations according to the ANSI S3.5. When using the MAF values according to ANSI S3.6 one has to calculate the CFFD with equation x